## Slide 1: EE93 – Medical Mobile Devices and Apps
Lecture: Instrumentation & DSP
60
98
30
bpm
%
rpm

## Slide 2: ECG Waveform on Strip Chart
1 mV, 10 mm high reference pulse
Length: 0.200 s
5 mm by 5 mm reference square
0,200 s duration by 0.5 mV amplitude
1 mm by 1 mm reference square
0,040 s duration by 0.1 mV amplitude
One heartbeat cycle
12-lead – showing in 4 columns by 3 rows

## Slide 3: Measuring ECG (3-Lead)
60
98
30
bpm
%
rpm
Source for ECG slides: Computing the Electrical Activity in the Heart: 1 (Monographs in Computational Science and Engineering) by Joakim Sundnes, Glenn Terje Lines, Xing Cai and Bjørn Frederik Nielsen (2007)
3-lead ECG uses right arm (or chest), left arm (or chest) and left foot
Able to obtain PQRST wave
Unable to obtain other leads and heart angle

## Slide 4: Common Frequencies for ECG
Heart rate: 0.67 to 5 Hz (40 to 300 bpm)
P-wave: 0.67 to 5 Hz
QRS Complex: 10 to 50 Hz
T-wave: 1 to 7 Hz
High frequency potentials: 100 to 500 Hz

## Slide 5: Common Frequencies for ECG Artifacts & Noise
Muscle: 5 Hz to 50 Hz
Respiratory: 0.12 to 0.5 Hz (8 to 30 bpm)
External Electric: 50 Hz or 60 Hz (AC Line)
Other Electrical: > 10 Hz (muscle stimulators, magnetic fields, pacemakers with impedance monitoring)

## Slide 6: ECG Special Notes
Skin-electrode interface – largest source of interference – produces 200 to 300 mV
Skin-electrode interference is magnified by motion (patient movement, respiratory variation)
Electrical activity of heart – 0.1 to 2 mV

## Slide 7: Power Spectra of ECG
Relative power spectra of QRS complex, P and T waves, muscle noise and motion artifacts based upon an average of 150 bpm
Source: http://www.ems12lead.com/wp-content/uploads/sites/42/2014/03/ecg-component-frequencies.jpg

## Slide 8: ECG Amplifier
V2
+
V1
+

## Slide 9: Signal & Noise Model
+
-
Vsignal + Vnoise
Vnoise
Vsignal

## Slide 10: Instrumentation Amplifier
V1
Vout
–
+
–
+
–
+
V2
R2
R1
R2
R3
R3
R4
R4

## Slide 11: Instrumentation Amplifier (IA)
Provides capability to:
  Reject common-mode signal components (noise & interference, undesired DC offsets)
  Amplifies differential-mode signal
In practice, rejection of common-mode signal is not complete  common-mode rejection ration (CMRR)

## Slide 12: Instrumentation Amplifier (IA)
Provides impedance isolation between bridge transducers and differential amplifier stage
Signals V1 and V2 are amplified separately
Conditions the signals
Provide high CMRR if implemented with diligence

## Slide 13: Instrumentation Amplifier
V1
Vout
–
+
–
+
–
+
V2
R2
R1
R2
R3
R3
R4
R4

## Slide 14: Level Shifter
Wide spread use in medical applications
Adds or subtracts a DC offset to or from signal
–
+
V+
Vout
RF
Rs
+
-
Vref

## Slide 15: Signal Processing
Instrumentation Amplifier
High Pass Filter
Pulse Indicator
Signal Processing
WiFi
Patient Monitor
ECG with Noise
Stop Band
Filter
Square Signal
Pulse Detect

## Slide 16: DSP
 IIR Filter
 FIR Filter

## Slide 17: Filter Specification
“Effective edge of the filter”
“Ripple”
“Ripple”

## Slide 18: DSP Notes
IIR filter – has infinite impulse response  need to limit
FIR filter – has finite impulse response  hf[n] = 0, n ≥ 0
FIR filter advantages:
  Can have exact linear phase
  Always stable (even under quantization)
  Design methods are reasonable linear
  Realize efficiently in hardware or software
  Transients have finite duration
Disadvantages
  Requires higher filter order that IIR to achieve similar performance
  Delay is typically greater in FIR than IIR counterpart

## Slide 19: FIR Filter Design Notes
IIR: H[Ω] = desired IIR filter with impulse h[n]
FIR:
Transfer function:
DTFT:

## Slide 20: DSP – Analytically
hd[n] = w[n]h[n]
  Where w[n] is a window function  truncates the signal
  Rectangular window causes abrupt transitions
  Other windows allow gradual transitions

## Slide 21: DSP – Other Windows
Hanning:
Hamming:

## Slide 22: DSP – Windows
Hd(Ω) better approximates H(Ω) when main lobe of filter is narrow in bandwidth and side lobes are small in value
Hanning and Hamming, in general have much smaller sidelobes than rectangular window  less ripple in frequency response of FIR filter

## Slide 23: DSP – Procedure
 signal that needs to be filtered
  Design the filter
  Normalize the Nyquist rate across the spectrum
  Generate the filter coefficients in MatLab
    Use MatLab command fir1
  Iterate until you “get an acceptable response”
    Use MatLab command filter on signal
 signal filter in iPad
  Set up difference equation
  Use filter coefficients from fir1
  Compute filtered signal in code using add/multiply via difference equation
  Program filter in Objective-C – rather than vDSP framework